## Slide 1: FLC-PCR Update 5-28-2020
James Ferguson

## Slide 2: FLC-PCR Software Updates
Fixed a step timing issue
Added data logging for graphing
Updated old Debug UI and logger to somewhat work for PID tuning with the new software
Added NAN checking and avoidance for temperature
Added folders to file save load

## Slide 3: Acquired FLC-PCR Hardware and Setup

## Slide 4: The Old PCB is Broken
Serial Communication Stability Problems
Board no longer current driver signals grounded seemed to randomly go low
Moved to bread bord temporarily and noted problems disappear

## Slide 5: Sensor Filtration Revisited
PSU on Hardware Filter
PSU on Software Filter
PSU off
PSU on
1.59Hz Low Pass Filter 10kΩ & 10uF
2 weighted average of the current and last value filters at Sensor In and PWM Out

## Slide 6: FLC-PCR PCB
Made with Autodesk EAGLE
38.72mm x 56.50mm
2 layer, all THP component design
5 PCB’s where ordered from PCBWay costing a total of $24.00
rev 1

## Slide 7: PCBs Arrived Today!
Checked for continuity, Matches Design.
Silkscreen Font was not consistent with the design for non vectorized text.
Part values where not printed.

## Slide 8: Bill of Materials
Item | Amount | Unit Price | Image
100nF Capacitor | 2 | $0.35 | 
16MHz Resonator | 1 | $0.95 | 
DIP Sockets Solder Tail - 28-Pin | 1 | $0.95 | 
Break Away Male Headers - Right Angle | 1 | $1.95 | 
Header - 10-pin Female | 2 | $0.50 | 
Break Away Male Headers – Straight | 1 | $1.50 | 
6mm Push Button | 1 | $- | 
10kΩ Resistor | 3 | $- | 
1kΩ Resistor | 1 | $- | 
1MΩ Resistor | 1 | $- | 
5mm LED | 1 | $- | 
ATmega328P | 1 | $- | 

## Slide 9: Thermal Performance Issue
Feb 13, 2020
May 22, 2020
The May 22, 2020 never changes temperature because the system was unable to get close enough to the target temperature to trigger the next step.
Temperatures are estimates of PCR tube temperature

## Slide 10: Possible Causes
Software problems
  Ran both tests using the same software, settings and experiment Feb 13, 2020 commit 4a1461718987835f0 d123c65a39d32330f6f26af.
Thermal Paste Failure
  Didn't think of this until after manufacturing the 3rd attempt.
Hardware Failure
  Power supply still provides 15V.
  Current drivers still output 15V at max PWM.
  Peltier MTBF 17,500 cycles [1]. Upper estimated at 8575 cycle (245 days @ 35 cycles).
  2% increase in electrical resistance before failure [1].
Ambient temperature
  It was about 20ºC when I ran the May 22, 2020 test and I remember it being warmer on Feb 13, 2020.
[1] “Thermoelectric Module Reliability,” Thermoelectric. [Online]. Available: https://thermal.ferrotec.com/technology/thermoelectric-reference-guide/thermalref10/. [Accessed: 26-May-2020].

## Slide 11: Manufacturing
Reflectix Spiral Pipe Wrap 2”
PRUSA i3 MK2
HATCHBOX ABS White 1.75mm
Autodesk Fusion 360
Slic3r PE
2 different printing methods
  Standard
  Concentric Infill only
    5 – 10 times Faster
    Somewhat like vase mode but allows for multiple shells
    Uses less plastic
    Additional design requirements
    Harder to print
Materials and Methods

## Slide 12: Insulated Thermal Box
Printed using concentric infill only.
2 0.4mm vertical shells with Reflectix in-between.
Only surrounds the sides of the tubes.
Rests on top of the block and beneath the lid.
No observable effect on thermal performance.
Attempt 1

## Slide 13: Insulated Thermal Box
Printed using concentric infill only.
Larger to close gaps with special holes for cable routing.
Failed due to excessive warping during part cooling of long edge walls.
Attempt 2

## Slide 14: Insulated Thermal Box
Standard printing methods.
Increased wall thickness to 3mm.
While running the top became warm and the sides remained at ambient.
There is a sub millimeter gap between the interior and the heated block.
The box sits at an angle due to the wires hitting the heatsink.
This design slightly improved thermal performance, but not enough to cycle with settings.
Attempt 3
Original Thermal Paste | New Thermal Paste | New Paste with Box
82ºC | 84ºC | 88ºC
Max temperature reached.
Temperatures are estimates of PCR tube temperature

## Slide 15: Pre Heating
Another possible cause might be the heatsinks lower temperature.
12711-9L31-09CQ peltiers have a DTmax of 68ºC.
Quick Preheat was performed by running the peltiers at max cooling for one minute.
Slow Preheat was performed by oscillating between 60ºC and 80ºC for five minutes.
Original Thermal Paste | New Thermal Paste | New Paste with Box | Quick Preheat | Slow Preheat
82ºC | 84ºC | 88ºC | 90ºC | 91ºC
Max temperature reached.
Temperatures are estimates of PCR tube temperature

## Slide 16: Thank you!
Questions?